Re-tender CUI/Atd/37/Civil Engg (2020-2021)

Computer Control Electro-hydraulic Servo

Universal Testing Machine with standard accessories

Computer Control Electro-Hydraulic Servo UTM with 3000kN Compression Frame (Or Equivalent)

Specifications:

² Sampling rate: 1000Hz

² Communication Processor, 100MHz

² Load resolution 180,000 steps, two ranges 2mV/V and 4mV/V

² RS485 Interface for external keyboard/Display supports only one device

² Drive Interface ±10V (16bit) analogue command output OR, digital command output

² Safety functions like 4DI

² Two Option slots for second analogue channel for strain gauge extensometer OR, two encoder channels for extensometer with one or two encoder.

² PC communication via USB or Ethernet

² Internal socket for serial Sensors (COM 1)

² Internal socket for Debug (COM 2)

Technical specification

1000H Frame

Max. Load(kN) 1000

Load measuring range 0.4~100%F.S.

Load accuracy (%) ±0.5

Deformation measuring range 0.4~100%F.S

Deformation accuracy (%) ±0.5

Displacement resolution (mm) 0.01mm (0.001mm with advanced encoder)

Test loading speed(mm/min) for WAW-1000H 0.5-50 (0.1-50 based on Doli controller and Moog valve)

Max. Crosshead moving speed (mm/min) 200

Stress control range 1~60(N/mm2)S-1

Strain control range 0.00007/s~0.0067/s

Tensile space(mm) 1000

Compression space(mm) 850

Piston stroke(mm) 250

Column Distance(mm) 590

Column Diameter(mm) 100

Working table size(mm) 670x880

Flat jaw (mm) 0-30, 10-40

Round jaw(mm) Φ9-20, 20-40, 40-60

Jaw length(mm) 110

Jaw width(mm) 110

Platen size(mm) Dia. 220x40

Bending span(mm) 50-500

Roller diameter (mm) Ф50

Roller length (mm) 160

Bending depth (mm) 180

Net weight (kg) 3500

Max. height(mm) 2800

Dimension of load frame3(mm) 980X670X2700

Size of power pack(mm) 700x750x1100

Oil tank volume(L) 110

Oil pressure (MPa) 26

Power supply 3PH, 380VAC, 50H, 4kW

Standard accessories

1. Hydraulic tensile fixture for round sample dia. 9-20, 20-40, 40-60mm, flat sample: thickness 0-30, 10-40mm

2. Flexure/Bending fixture 1 set

Bending span: 50-500mm, roller dia. 50mm, upper punching head/mandrel: dia. 50mm,

3.Compression fixture dia.220x40mm

4.Computer system with English version software (19in LCD)

Computer: Dell, Intel (R) Celeron (R) 2.6Ghz, 2MB processor, 2GB, DDR3 1600MHz, SDRAM Memory, HD 250G,

16X DVD-ROM Drive, Windows 7 home basic system; 19inch screen LED monitor, photoelectrical mouse, and

Hydraulic shrinkage mould with measurement apparatus

multiple keyboard; Other advanced configuration can be offered upon request, but above configuration are quite enough for free operating of such machine.

3000kN Compression frame

Max. Load (kN) 3000

Load measuring range 0.4-100% of F.S.

Load accuracy Class 1

Constant loading rate 1-50kN/s (can be adjusted)

Back to zero relative error ±1%

Relative resolution 1%

Test space adjusting mode By dual action actuator & spacers

Compression platen (mm) Φ300mm

Distance between two platens(mm) 400mm

Distance between columns 350x350mm

Ram travel 250mm

Max. moving speed of piston 100mm/min

Dimension for load frame 630x630x1450mm

Weight 2800kg

Options:

Full-enclosed safety cover with metal sheet around the load

frame of UTM with electromagnetic switch for protection

Air cooling system to cool down the hydraulic oil if for long term operation

Servo valve from Moog, USA

Load cell from Celtron, USA

Advanced encoder with 40000pulse/turn can achieve better

position resolution 0.001mm

clip on axial extensometer from Epsilon, USA

with 50mm gauge length, -5~+10mm travel, for flat

and round samples, accuracy: class 0.5 as per ISO 9513

S p e c i f i c a t i o n s

Excitation: 5 to 10 VDC recommended, 12 VDC or VAC max.

Output: 2 to 4 mV/V, nominal, depending on model

Linearity: ≤0.15% of full scale measuring range

Temperature Range: Standard (-ST) is -40 °C to +100 °C (-40 °F to 210 °F)

Optional (-LHT) is -270 °C to +200 °C (-454°F to 400 °F)

Cable: Integral, ultra-flexible cable, 2.5 m (8 ft) standard

Standard Quick Attach Kit: Fits round samples up to 25 mm diameter (1.0 inch) and flats to

12 mm thick by 31 mm wide (0.5 inch by 1.25 inch)

Operating Force: 30 g typical

EDC 220 Controller from Doli, Germany with self-recognition and

auto-calibration function for both load cell and extensometer

Industrial Touch Screen Pad 24inch

Shear test fixture for dia. 10mm sample

Shear test fixture for concrete dia. 50mm, dia. 100mm

Shear test fixture for bolt M6-M30

Special tensile grip for fasteners as per ASTM F606 & ISO 898: M6, M7, M8, M10, M12, M14, M16, M18, M20, M22, M24, M27, M30

Special grips for pre-stressed 7-wire strand dia. 9.5, 12.7,

15.2mm as per ASTM A416 with high pressure power pack

CMOD testing Fixture as per EN 14651

COD gage for determination of fracture toughness properties of materials

S P E C I F I C AT I O N S

Excitation: 5 to 10 VDC recommended, 12 VDC or VAC max.

Output: 2 to 4 mV/V nominal, depending on model

Linearity: 0.15% for measuring ranges less than 0.25 inch

(6 mm), 0.20% for greater measuring ranges

Temperature Range: Standard (-ST) is -40 °C to +100 °C (-40 °F to 210 °F)

Cable: Integral, ultra-flexible cable, 8 ft (2.5 m) standard

Operating Force: Exerts 2 to 3 lbs (9 to 14 N), depending on model

Axial Averaging Extensometer for Modulus of Elasticity, and together with model 3544-200M- 060M-ST Circumferential Extensometer for Poisson Ratio Determination

Specifications

Excitation: 5 to 10 VDC recommended, 12 VDC or VAC max.

Output: 2 to 4 mV/V nominal, depending on model

Linearity: ≤0.20% of full scale measuring range, depending

on model

Temperature Range: Standard (-ST) is -40 °C to +100 °C (-40 °F to 210 °F)

Optional (-LHT) is -270 °C to +200 °C (-454°F to 400 °F)

Cable: Integral, ultra-flexible cable, 2.5 m (8 ft) standard

Specimen Size: Wide range of specimen sizes available, including AX, BX, NX and

larger, to 200 mm (and 8 inches) diameters standard (larger sizes on special order)

Operating Force: <30 g typical per side

RTTS-100D Dual Camera 3D DIC Video extensometer for axial & transverse strain, Modulus of Elasticity, Poisson ratio

measurement etc. for all metals and concretes

Specifications:

Frequency @ FOV Class 1 according to ISO 9513/

Frequency @ FOV Class A according to ASTM E83

60Hz @ 100x50mm

Classification to ISO 9513/ASTM E 83 Class 0.5/Class B-1

Field of View (FOV) 1-100mm

Gauge Range 0.5-100 % FOV

Resolution

0.19 μm @ 53 mm FOV

0.39 μm @ 100 mm FOV

Digital output RS232/485, TCP/IP

Camera resolution 5 MPx @9fps (Max. 200fps AOI)

Electro-hydraulic Servo Dynamic and Static

Universal Testing Machine with standard accessories

Main specifications:

1. Max. Static load: ±100kN

2. Max. Dynamic load: ±70kN

3. Max. Stroke of the actuator: ±75mm with accuracy from less than ±0.5 % of full scale, total 150mm

4. Load accuracy: ±1%

5. Deformation measuring accuracy: ±1%

6. Frequency range: 0.1-50Hz with EDC580 controller;

7. Main wave shapes: sine wave, square wave, triangular wave, trapezoidal, static test oblique wave.

8. Actuator Diameter of rod/cylinder: Φ70/110mm

9. Horizontal Clearance: 600mm

10. Vertical clearance: 1300mm/extended space: 1800mm

11. Working height between grips: 950mm

12. Specimen clamping: Hydraulic forcing clamping.

13. Crosshead moving: hydraulic driving for elevating.

14. Stiffness of load frame: 450kN/mm

15. Constant pressure servo pump station: Flow 45L/min, pressure 210Bar, Motor power 22kW, 380V, 50Hz

16. Controller: Doli EDC 580 full-digital intelligent controller

17. Size: load frame: 1000x650x2900/3400mm (standard/extended); hydraulic power pack: 1300x1030x1500mm

18. Weight of load frame: 1500kg; Weight of hydraulic power pack: 500kg (Or Equivalent)